## Sheet1
Routine Test Certificate
Tested Product
Commercial Reference : | [10 empty] | Serial Number :
Test Condition
Test Date: | [10 empty] | Reference Meter: |  | CLOU CL311V2
Temperature: |  |  |  | 23℃ ± 2 ℃ | [6 empty] | Humidity: | 50% ± 10%
Injected values :
| Item |  | Ia(A) | Ib(A) | Ic(A) | Va(V) | Vb(V) | Vc(V) | PF | F(Hz) | Iref = current reference, designation according to
| 1 | Imin | 0.5 | 0.5 | 0.5 | 230 | 230 | 230 | 1 | 50 | standard EN50470-1
| 2 | 10Itr | 10 | 10 | 10 | 230 | 230 | 230 | 1 | 50
| 3 | 10Itr | 10 | 10 | 10 | 230 | 230 | 230 | 0.5L | 50 | In = value of current reference (n=phase a to c)
| 4 | Imax | 63 | 63 | 63 | 230 | 230 | 230 | 1 | 50
| 5 | 10Itr | 10 | 0 | 0 | 230 | 230 | 230 | 1 | 50 | Vn = value of voltage reference (n=phase a to c)
| 6 | 10Itr | 0 | 10 | 0 | 230 | 230 | 230 | 1 | 50
F = frequency reference
PF = power factor
Measured Results
| Item | MPE(%) | Ic(A): Limit -; Vc(V): Limit + | MPE(%) = measured error, defined by the formula :
| 1 |  | Ic(A): 0; Vc(V): 2
| 2 |  | Ic(A): 0; Vc(V): 2 | MPE(%) = [(Emes – Eref)/Eref]*100%
| 3 |  | Ic(A): 0; Vc(V): 2
| 4 |  | Ic(A): 0; Vc(V): 2 | (Emes = measured energy, Eref = energy reference)
| 5 |  | Ic(A): 0; Vc(V): 2
| 6 |  | Ic(A): 0; Vc(V): 2 | Limit + = Maximum error tolerated defined by
standard EN50470-3 class B;
The counter is declared compliant if for each Iref:
Limit- <=MPE <= Limit+
Conformity Approval
This calibration certificate guarantees that the results
comply with the international unit system(IS) | [10 empty] | compulsory approval to validate the certificate
Zhou Lin